## Cover-Page
|  |  | Title: | VERASENSE FTS Support
|  |  | Number: | DWG-01000-001
|  |  | Revision: | 11
NOTES
1. Do not modify this file without Engineering approval. 2. Copied and renamed versions of this file may be used for engineering evaluation of unit performance. 3. Device type information on "FTS-Initialization" must be updated for new VERASENSE products. 4. Spreadsheet currently configured only for the following device types. Do not use for production of other devices without engineering approval: · SYK 1204 · BMT 1204 · ZMH 1204 · SNN JRNY 1204 · SNN LGN 1204
REVISION HISTORY
REV | development rev. | DCO# | DESCRIPTION |  | INITIATOR | DATE
A |  | - | Unreleased |  | Andrew Miller | 2014-01-06
B |  | - | Initial Release; Adds Min/Max Failure Modes at 1204 Scale Calculation |  | Andrew Miller | 2014-01-08
C |  | - | Changed Scale Calc to 10-40; increased dwell time; better error code report |  | Andrew Miller | 2014-01-13
D |  | - | BMT accelerometer default value fix |  | Andrew Miller | 2014-01-17
E |  | - | Max Scale Factor; More error codes; Supports 51 force points; 41k Batt Minimum; Angles Defined |  | Andrew Miller | 2014-01-22
4 | F | - | Fixed Notes on Cover Page, Rounded Accel Values Safely; 20Hz Transmission (GCLKV->BB) |  | Andrew Miller | 2014-02-06
5 | G |  | Increased 1204 Batt Max limit to 52000. Decreased 1204 Batt Min limit to 40000 |  | Juvencio Dieguez | 2014-02-20
6 | H |  | Cleared all fields of pre-existing readings, except placed default accelerometer data and default Lot and serial number on 1204-Initial. Changed GainC to Operative Side on E-44 at the 1204Scale-Volt-acc tab and linked to 1204-Initial F-4 for SNN leg id (76=L,82=R, 255 is default) on 1204 tabs. Default Dev Type is 1013. Added Smith&Nephew codes in the 7000's section. In 1204-Scale-Volt-Accel added 2.8V section at J31, revised L21 to be both Z or 2.8V, added P72 new load section battery and battery/scale areas at K169 through Q169. Changed Batt Min limit to 30000. All GAINC values look at the GAINC value on the FTS-Initialization page. |  | Tim Labuhn | 2014-08-01
7 | I |  | Revert back the changes done on 2/6/14 to EEPROM address 29 for the 20Hz Transmission (GCLKV) from BB back to BF. |  | Juvencio Dieguez | 2014-10-10
8 | J |  | Revised Error calc in C137 and P72 to S79 and added table U78 to U81 - Page 1204-Scale-Volt-Accel. Changed colors of K10 and L10. Changed error fail code in Q171. Added R167 to R171. Changed D36 to D43, added error code 43; Changed default serial number to 13466634 on 1204-Initial; Added PASS/FAIL and error code display next to 1204-Results on Presentation page and also added Cal Factor display. Replaced Depuy with SNN Legion 8000 catalog numbers. Changed SNN prefix to SNN-JRNY and SNN-LGN. |  | Juvencio Dieguez | 2014-10-17
9 | K |  | Changed error 40 to errors 72 to 77 defining which sensor caused the error, add table of increasing loads error check adding errors 66 to 71 (these codes are not included as failures). Changed cells 83B,C and D to correct math error and calibrate between 0 and 80. Add Count/LB at B62 to H62, Added scale multiplier at E84 and F84, Revised the load calculations from C74 to H79, C87 to J92 and C96 to D101, Revised failure code cell C137, Changed Accelerometer error codes from 70 - 75 to 44 - 49 to make room for sensor errors. On FTS-Initialization AE21 was 20, AE32 was 120, AE28 was 65000, AE29 was 1000 (these new values are based on past tests, not statistics). Made B28 = 5 dwell, was 10. Changed error fail code in Q171 to cover also the Count Min/Max Test. |  | Tim Labuhn | 2015-04-24
10 | L |  | Added Smith & Nephew size 1-2 for Journey and Legion families. Moved the Catalog Number Lookup Table around to accommodate the longer SNN-JRNY table and split the BCS from the CR. On FTS-Initialization changed RF transmit power setting PATABLE from 77 to 84 on all OEMs except BMT that was already at 84. Made B26=10 dwell, was 5. |  | Juvencio Dieguez | 2015-06-23
11 | M |  | a)-In 1204-Scale-Volt-Accel, on E83 and F83 cells, "condition" Scale Factor to be a value different from 0 to avoid "DIV/0" erros and false Pass devices. Made M169 and N169 same as Scale Factor A and Scale Factor B. b)- created date and time stamp slot on Presentation tab: AC3=Date AD3=Time. c)-Added Error code 80 through 86 to cover Load Accuracy test fail errors. d)-Added table on the FTS-Initialization tab for Load Accuracy Force loads and dwell time. e)-on 1204-Scale-Volt-Accel tab, added tables (no yet used) to capture Load Acc test raw data, filtered data and average data. Added table for calculated load per sensor and for errors. f)-Added logic for P/F criteria and error code. g)-Mirrowed the Load Accuracy Force table with added error columns on the Presentation tab and inserted plot chart for visualization. h)-Created place holder on 1204-Scale-Volt -Accel tab: O207=Correction Factor. i)-condition the logic of error 86 to display 0 if any other error was displayed, done to allow the test to continue. j)-increased the Batt Min counts from 30000 to 33000, apx 2.8V. k)-on 1204-Scale-Volt-Accel tab added L171 to capture error based on Batt A/D count at the end of the test. Made L171 and P171 part of sum for C137. l)-on 1204-Scale-Volt-Accel tab changed V4 (3.1V) on K27 to look at L169 instead of J169. m)-on 1204-Scale-Volt-Accel tab changed table "New Loads converted" to use scale factors on E137 and G137 instead of values from EEprom on M169 and N169. Added H to S variable on ESM Op table at Initialization tab. n)-on 1204-Scale-Volt-Accel tab changed formula on J158 to account for 0 degree and 90 degree incorrect placement at 45degree angle capture. o)-on 1204-Scale-Volt-Accel tab, cell C137 changed the logic to condition the error shown to display Accel or Load Test or Batt error related instead of overall Load Accuracy. p)-On 1204-Scale-Volt-Accel tab cell C137, took P171 off the logic for errors. q)-On 1204-Scale-Volt-Accel tab cell L101, added Program Error box. r)-Load Accuracy error table is now +/-3.5, before it was only +3.5. H151 default changed to 48000. |  | Juvencio Dieguez | 2016-05-24
12 | N | - | a)-In 1204-Scale-Volt-Accel, on cellsS81, U81,V81,W81,X81,Y81,Z81 conditioned the cells to display 0. The error codes displayed previously on these cells are now shifted and displayed to the cells below: S82,U82,V82,W82,X82,Y82,Z82. |  | Juvencio Dieguez | 2016-06-16
13 | O |  | Includes new speed configurations. |  | Wael Hazin | 2016-07-15

## Presentation
[15 empty] | Error Code | [8 empty] | Time Stamp
[14 empty] | FTS only |  | Voltage Tool |  |  |  | ScaleA | ScaleB | Error? |  | Date | Time
1010 Results | [10 empty] | 1204 Results |  |  | FAIL | 77 | 43 | FAIL |  |  | -150 | -150 | PASS | FAIL | 2016-07-16 | 22:34
F/Cond [lb] | S3 [lb] | S4 [lb] | S5 [lb] | S6 [lb] | S7 [lb] | S8 [lb] | CondA [lb] | CondB [lb] | ErrorA [lb] | ErrorB [lb] | F/Cond [lb] | S3 [lb] | S4 [lb] | S5 [lb] | S6 [lb] | S7 [lb] | S8 [lb] | CondA [lb] | CondB [lb] | ErrorA [lb] | ErrorB [lb] | ErrorMax | ErrorMin
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.5 | -3.5
5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -5 | -5 | 5 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 35 | 35 | 3.5 | -3.5
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -10 | -10 | 10 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 30 | 30 | 3.5 | -3.5
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -20 | -20 | 20 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 20 | 20 | 3.5 | -3.5
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -30 | -30 | 30 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 10 | 10 | 3.5 | -3.5
40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -40 | -40 | 40 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 0 | 0 | 3.5 | -3.5
Load Accuracy Test
[11 empty] | F/Cond [lb] | S3 [lb] | S4 [lb] | S5 [lb] | S6 [lb] | S7 [lb] | S8 [lb] | CondA [lb] | CondB [lb] | ErrorA [lb] | ErrorB [lb] | ErrorMax | ErrorMin
[11 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.5 | -3.5
[11 empty] | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -10 | -10 | 3.5 | -3.5
[11 empty] | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -20 | -20 | 3.5 | -3.5
[11 empty] | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -40 | -40 | 3.5 | -3.5

## FTS-Initialization
Catalog Number Lookup Table
UDF1 |  | 1000 | UDF2 |  | 2000 | SYK |  | 3000 | BMT |  | 4000 | MKO |  | 5000 | ZMR | 6000 | SNN-JRNYBCS |  | 7000 | SNN-JRNYCR |  | 7000 | SNN-LGN |  | 8000 | UDF3 |  | 9000
Undefined |  | 1001 | Undefined |  | 2001 | SYK-TRCR01 |  | 3001 | BMT-VGCR63 |  | 4001 | MKO-MCK04 |  | 5001 | ZMR-NGCRCH-34 | 6001 | SNN-JRNYBCS12-L |  | 7009 | SNN-JRNYCR12-L |  | 7017 | SNN-LGNPS12 |  | 8001 | Undefined |  | 9001
Undefined |  | 1002 | Undefined |  | 2002 | SYK-TRCR02 |  | 3002 | BMT-VGCR71 |  | 4002 | Undefined |  | 5002 | ZMR-NGCRCH-56 | 6002 | SNN-JRNYBCS12-R |  | 7010 | SNN-JRNYCR12-R |  | 7018 | SNN-LGNPS34 |  | 8002 | Undefined |  | 9002
Undefined |  | 1003 | Undefined |  | 2003 | SYK-TRCR03 |  | 3003 | BMT-VGCR79 |  | 4003 | Undefined |  | 5003 | ZMR-NGCRCH-70 | 6003 | SNN-JRNYBCS34-L |  | 7011 | SNN-JRNYCR34-L |  | 7019 | SNN-LGNPS56 |  | 8003 | Undefined |  | 9003
Undefined |  | 1004 | Undefined |  | 2004 | SYK-TRCR04 |  | 3004 | BMT-VGPS63 |  | 4004 | Undefined |  | 5004 | ZMR-NGPSCD-34 | 6004 | SNN-JRNYBCS34-R |  | 7012 | SNN-JRNYCR34-R |  | 7020 | SNN-LGNPS78 |  | 8004 | Undefined |  | 9004
Undefined |  | 1005 | Undefined |  | 2005 | SYK-TRCR05 |  | 3005 | BMT-VGPS71 |  | 4005 | Undefined |  | 5005 | ZMR-NGPSEF-34 | 6005 | SNN-JRNYBCS56-L |  | 7013 | SNN-JRNYCR56-L |  | 7021 | SNN-LGNCR12 |  | 8005 | Undefined |  | 9005
Undefined |  | 1006 | Undefined |  | 2006 | SYK-TRCR06 |  | 3006 | BMT-VGPS79 |  | 4006 | Undefined |  | 5006 | ZMR-NGPSEF-56 | 6006 | SNN-JRNYBCS56-R |  | 7014 | SNN-JRNYCR56-R |  | 7022 | SNN-LGNCR34 |  | 8006 | Undefined |  | 9006
Undefined |  | 1007 | Undefined |  | 2007 | SYK-TRCR07 |  | 3007 | Undefined |  | 4007 | Undefined |  | 5007 | Undefined | 6007 | SNN-JRNYBCS78-L |  | 7015 | SNN-JRNYCR78-L |  | 7023 | SNN-LGNCR56 |  | 8007 | Undefined |  | 9007
Undefined |  | 1008 | Undefined |  | 2008 | Undefined |  | 3008 | Undefined |  | 4008 | Undefined |  | 5008 | Undefined | 6008 | SNN-JRNYBCS78-R |  | 7016 | SNN-JRNYCR78-R |  | 7024 | SNN-LGNCR78 |  | 8008 | Undefined |  | 9008
Undefined |  | 1009 | Undefined |  | 2009 | Undefined |  | 3009 | Undefined |  | 4009 | Undefined |  | 5009 | Undefined | 6009 | Undefined |  | 7030 | Undefined |  | 7025 | Undefined |  | 8009 | Undefined |  | 9009
Undefined |  | 1010 | Undefined |  | 2010 | Undefined |  | 3010 | Undefined |  | 4010 | Undefined |  | 5010 | Undefined | 6010 | Undefined |  | 7031 | Undefined |  | 7026 | Undefined |  | 8010 | Undefined |  | 9010
Undefined |  | 1011 | Undefined |  | 2011 | Undefined |  | 3011 | Undefined |  | 4011 | Undefined |  | 5011 | Undefined | 6011 | Undefined |  | 7032 | Undefined |  | 7027 | Undefined |  | 8011 | Undefined |  | 9011
Undefined |  | 1012 | Undefined |  | 2012 | Undefined |  | 3012 | Undefined |  | 4012 | Undefined |  | 5012 | Undefined | 6012 | Undefined |  | 7033 | Undefined |  | 7028 | Undefined |  | 8012 | Undefined |  | 9012
Undefined |  | 1013 | Undefined |  | 2013 | Undefined |  | 3013 | Undefined |  | 4013 | Undefined |  | 5013 | Undefined | 6013 | Undefined |  | 7034 | Undefined |  | 7029 | Undefined |  | 8013 | Undefined |  | 9013
Force [lb] Targets per Type. Controls Application of FTS loads per device type | [11 empty] | Pass/Fail Criteria, 1010 |  |  |  | OEM Lookup Table | [6 empty] | ESM301 Operation |  |  |  | Pass/Fail Criteria, 1204
Dwell [s] | Default | 5000 | Unused | Unused | Unused | Unused | Unused | Unused | Unused | Unused |  | Limit | Value | Units | Toggle | OEM |  | Prefix | Series | PATABL | UNI/TKA |  | Param. | Value | Units |  | Limit | Value | Units | Toggle
[12 empty] | Ambitious | 3.5 | lbs | TRUE | Undefined |  | UDF1 | 1000 | 84 | TKA |  | Kp | 0.03 | - |  | Ambitious | 3.5 | lbs | TRUE
0.1 | 70 | 35 | [9 empty] | Min Scale | 40 | - | TRUE | Undefined |  | UDF2 | 2000 | 84 | TKA |  | Ki | 2e-05 | - |  | Min Scale | 30 | - | TRUE
0.1 | 140 | 70 | [13 empty] | Stryker |  | SYK | 3000 | 84 | TKA |  | Kd | 0.002001 | - |  | 0g Min | 22000 | - | TRUE
0.1 | 0 | 0 | [13 empty] | Biomet |  | BMT | 4000 | 84 | TKA |  | Force Tol. | 1 | lb |  | 0g Max | 42000 | - | TRUE
0.1 | 70 | 35 | [13 empty] | Mako |  | MKO | 5000 | 84 | UNI |  | H to S | 2 | in |  | 1g Min | 45000 | - | TRUE
0.1 | 140 | 70 | [13 empty] | Zimmer |  | ZMR | 6000 | 84 | TKA | [5 empty] | 1g Max | 65000 | - | TRUE
10 | 0 | 0 | [9 empty] | BattMax | 24000 | - | TRUE | Smith & Nephew Journey BCS |  | SNN-JRNYBCS | 7000 | 84 | TKA |  | Default Values |  |  |  | BattMax | 52000 | - | TRUE
5 | 10 | 5 | [9 empty] | BattMin | 18000 | - | TRUE | Smith & Nephew Journey CR |  | SNN-JRNYCR | 7000 | 84 | TKA |  | Param. | Value | Units |  | BattMin | 33000 | - | TRUE
5 | 20 | 10 | [9 empty] | CntMax | 65000 | - | TRUE | Smith & Nephew Legion |  | SNN-LGN | 8000 | 84 | TKA |  | 1010GAINC | BF | - |  | CntMax | 60000 | - | TRUE
5 | 40 | 20 | [9 empty] | CntMin | 40000 | - | TRUE | Undefined |  | UDF3 | 9000 | 84 | TKA |  | Scale1010 | 100 | - |  | CntMin | 5000 | - | TRUE
5 | 60 | 30 | [9 empty] | Anomaly | 30 | % | TRUE | Undefined |  | UDF4 | 10000 | 84 | TKA |  | Scale1204 | 100 | - |  | Anomaly | 30 | % | FALSE
10 | 80 | 40 | [9 empty] | Negative | -1 | - | TRUE | Undefined |  | UDF5 | 11000 | 84 | TKA |  | Voltage | 0 | - |  | Negative | -1 | - | FALSE
0 | -1 | -0.5 | [12 empty] | FALSE | Undefined |  | UDF6 | 12000 | 84 | TKA |  | 1204GAINC | EF | - |  | MaxScale | 70 | - | TRUE
FALSE
[12 empty] | Error Code Table | [13 empty] | Sensor Channels
[12 empty] | Code | Stage | Topic |  |  | Description | [8 empty] | Sensor | UNI | TKA
[12 empty] | 0 | All | No Error |  |  | No Error Reported; Test Continuing | [8 empty] | 0 | 2 | 2
[12 empty] | 1 | Startup | FTS Application Error |  |  | Software Installation Issue | [8 empty] | 1 | 2 | 2
[12 empty] | 2 | Startup | FTS Application Error |  |  | Transceiver Communication Error | [8 empty] | 2 | 2 | 2
[12 empty] | 3 | Startup | FTS Application Error |  |  | Mark-10 Communication Error | [8 empty] | 3 | 4 | 4
[12 empty] | 4 | Startup | FTS Application Error |  |  | Missing or Bad Support File | [8 empty] | 4 | 4 | 4
[12 empty] | 5 | Startup | FTS Application Error |  |  | Can't Write to Output File | [8 empty] | 5 | 4 | 4
[12 empty] | 6 | Startup | Labview file Error |  |  | The FTS software could not write the data to the excel script | [8 empty] | 6 | 4 | 4
[12 empty] | 7 | [13 empty] | 7 | 0 | 4
[12 empty] | 8 | [13 empty] | 8 | 0 | 4
[12 empty] | 9 | [13 empty] | 9 | 0 | 4
[12 empty] | 10 | Initial | Programming Error |  |  | Timeout without DUT response
[12 empty] | 11 | Initial | Programming Error |  |  | Timeout after some Communication
[12 empty] | 12 | Initial | Battery Voltage Error |  |  | Voltage too Low at Initial Program | [8 empty] | Home To Surface Speed
[12 empty] | 13 | Initial | Battery Voltage Error |  |  | Voltage too High at Initial Program
[12 empty] | 14 | Initial | PCB Error |  |  | One or more Sensors significantly different from Median Value | [8 empty] | Distance from Home (%) |  | Speed (in/min
[12 empty] | 15 | Initial | PCB Error |  |  | One or more Sensors exceed maximum Allowed Value
[12 empty] | 16 | Initial | PCB Error |  |  | One or more Sensors less than minimum Allowed Value | [8 empty] | 80 |  | 12 | within 0% to 79%
[12 empty] | 17 | Initial | Communication Error |  |  | Too many dropped Data Packets | [8 empty] | 85 |  | 8 | within 80% to 84%
[12 empty] | 18 | Initial | Communication Error |  |  | Device Under Test not transmitting on Expected Channel | [8 empty] | 90 |  | 6 | within 85% to 89%
[12 empty] | 19 | Initial | Labview file Error |  |  | The FTS software could not write the data to the excel script | [8 empty] | 95 |  | 4 | within 90% to 94%
[12 empty] | 20 | [13 empty] | 100 |  | 2 | within 95 to 100%
21
[12 empty] | 22 | [13 empty] | Mark-10 Home Speed |  |  | 35
[12 empty] | 23 | [13 empty] | Mark-10 Load Speed |  |  | -1 | Note: If load speed ≤ 0 the FTS will utilize the PID Controller
24
25
26
27
28
29
[12 empty] | 30 | Optim | Programming Error |  |  | Timeout without DUT response
[12 empty] | 31 | Optim | Programming Error |  |  | Timeout after some Communication
[12 empty] | 32 | Optim | Labview file Error |  |  | The FTS software could not write the data to the excel script
[12 empty] | 33 | Optim | Cannot Optimize Sensor |  |  | Sensor values exceed maximum allowable for optimization
[12 empty] | 34 | Optim | Cannot Optimize Sensor |  |  | Sensor values below minimum allowable for optimization
[12 empty] | 35 | Optim | Battery Voltage Error |  |  | Voltage too Low at Initial Program
Angle Targets per Type. Directs the User to Orient the VERASENSE for Accelerometer Calibration (NOT USED YET) | [11 empty] | 36 | Optim | Battery Voltage Error |  |  | Voltage too High at Initial Program
Dwell [s] | Orientation | User Text | [9 empty] | 37 | Optim | PCB Error |  |  | One or more Sensors significantly different from Median Value
2 | 1 | Position VERASENSE in Orientation A (Superior Side Up) and Select Capture. | [9 empty] | 38 | Optim | PCB Error |  |  | One or more Sensors exceed maximum Allowed Value
2 | 2 | Position VERASENSE in Orientation B (Anterior Side Up) and Select Capture. | [9 empty] | 39 | Optim | PCB Error |  |  | One or more Sensors less than minimum Allowed Value
2 | 3 | Position VERASENSE in Orientation C (Inferior Side Up) and Select Capture. | [9 empty] | 40 | Optim | Sensor Range Error |  |  | One or more Sensors less than minimum Allowed Value
2 | 4 | Position VERASENSE in Orientation D (Posterior Side Up) and Select Capture. | [9 empty] | 41
2 | 5 | Position VERASENSE in Orientation E (Left Side Up) and Select Capture. | [9 empty] | 42
2 | 6 | Position VERASENSE in Orientation F (Right Side Up) and Select Capture. | [9 empty] | 43 | Batt | Input Voltage |  |  | A/D Difference Outside Range
| -1 | [10 empty] | 44 | Angles | Accelerometer Error |  |  | Accelerometer X 0g value out of range
[12 empty] | 45 | Angles | Accelerometer Error |  |  | Accelerometer X 1g value out of range
[12 empty] | 46 | Angles | Accelerometer Error |  |  | Accelerometer Y 0g value out of range
[12 empty] | 47 | Angles | Accelerometer Error |  |  | Accelerometer Y 1g value out of range
[12 empty] | 48 | Angles | Accelerometer Error |  |  | Accelerometer Z 0g value out of range
[12 empty] | 49 | Angles | Accelerometer Error |  |  | Accelerometer Z 1g value out of range
[12 empty] | 50 | Forces | Communication Error |  |  | Too many dropped Data Packets
[12 empty] | 51 | Forces | Communication Error |  |  | Device Under Test not transmitting on Expected Channel
[12 empty] | 52 | Forces | Battery Voltage Error |  |  | Voltage too Low at Force Application
[12 empty] | 53 | Forces | Battery Voltage Error |  |  | Voltage too High at Force Application
Load Accuracy Force [lb] Targets per Type. Controls Application of FTS loads to device | [11 empty] | 54 | Forces | Accuracy Failure |  |  | Unknown Root Cause
Dwell [s] | Default | 5000 | Unused | Unused | Unused | Unused | Unused | Unused | Unused | Unused |  | 55 | Forces | Accuracy Failure |  |  | Likely bad zero Root Cause
[12 empty] | 56 | Forces | Accuracy Failure |  |  | Likely insufficient Preload Root Cause
10 | 0 | 0 | [9 empty] | 57 | Forces | PCB Error |  |  | Poor Sensor Sensitivity
10 | 20 | 10 | [9 empty] | 58 | Forces | Programming Error |  |  | Timeout without DUT response
10 | 40 | 20 | [9 empty] | 59 | Forces | Programming Error |  |  | Timeout after some Communication
10 | 80 | 40 | [9 empty] | 60 | Forces | Shutoff Error |  |  | Shutdown Timeout
0 | -1 | -0.5 | [9 empty] | 61 | Forces | Labview file Error |  |  | The FTS software could not write the data to the excel script
[12 empty] | 62 | Forces | PCB Error |  |  | One or more Sensors exceed maximum Allowed Value
[12 empty] | 63 | Forces | PCB Error |  |  | One or more Sensors less than minimum Allowed Value
[12 empty] | 64 | Forces | Accuracy Failure |  |  | Narrow Accuracy Failure. Unit may PASS a retest.
[12 empty] | 65 | Forces
[12 empty] | 66 | Forces | Loads are not increasing S3 |  |  | Sensor S3 value does not increase with increasing load
[12 empty] | 67 | Forces | Loads are not increasing S4 |  |  | Sensor S4 value does not increase with increasing load
[12 empty] | 68 | Forces | Loads are not increasing S5 |  |  | Sensor S5value does not increase with increasing load
[12 empty] | 69 | Forces | Loads are not increasing S6 |  |  | Sensor S6 value does not increase with increasing load
[12 empty] | 70 | Forces | Loads are not increasing S7 |  |  | Sensor S7 value does not increase with increasing load
[12 empty] | 71 | Forces | Loads are not increasing S8 |  |  | Sensor S8 value does not increase with increasing load
[12 empty] | 72 | Forces | Sensor S3 Range is too low |  |  | Sensor S3 range at 80lb is to low
[12 empty] | 73 | Forces | Sensor S4 Range is too low |  |  | Sensor S4 range at 80lb is to low
[12 empty] | 74 | Forces | Sensor S5 Range is too low |  |  | Sensor S5 range at 80lb is to low
[12 empty] | 75 | Forces | Sensor S6 Range is too low |  |  | Sensor S6 range at 80lb is to low
[12 empty] | 76 | Forces | Sensor S7 Range is too low |  |  | Sensor S7 range at 80lb is to low
[12 empty] | 77 | Forces | Sensor S8 Range is too low |  |  | Sensor S8 range at 80lb is to low
78
[12 empty] | 79 | Load Acc | Zero readings |  |  | Residual or Offset Load
[12 empty] | 80 | Load Acc | Compartment A outside spec |  |  | Compartment A at 10lb outside spec
[12 empty] | 81 | Load Acc | Compartment B outside spec |  |  | Compartment B at 10lb outside spec
[12 empty] | 82 | Load Acc | Compartment A outside spec |  |  | Compartment A at 40lb outside spec
[12 empty] | 83 | Load Acc | Compartment B outside spec |  |  | Compartment B at 40lb outside spec
[12 empty] | 84 | Load Acc | Compartment A outside spec |  |  | Compartment A at 80lb outside spec
[12 empty] | 85 | Load Acc | Compartment B outside spec |  |  | Compartment B at 80lb outside spec
[12 empty] | 86 | Load Acc | Overall Load Accuracy Test fail |  |  | Load Accuracy Test Incomplete or Failure
87
88
89

## EEPROM-Contents
ASIC 1010 EEPROM Contents |  |  |  | ASIC 1204 EEPROM Contents |  |  |  |  | ASIC 1010/1204 Mapping Contents Tables to EEPROM Programming Locations
Line | Address | Value | Description | Line | Address | Value | Description |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F
0 | 0x000 | 0x1B | Packet Size | 0 | 0x000 | 0x1B | Packet Size |  | 0x000 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
1 | 0x001 | 0x00 | SIVMK[7:0] | 1 | 0x001 | 0x00 | SIVMK[7:0] |  | 0x010 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
2 | 0x002 | 0x00 | SIVMK[15:8] | 2 | 0x002 | 0x00 | SIVMK[15:8] |  | 0x020 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47
3 | 0x003 | 0x00 | SULMK[7:0] | 3 | 0x003 | 0x00 | SULMK[7:0] |  | 0x030 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63
4 | 0x004 | 0x00 | SULMK[15:8] | 4 | 0x004 | 0x00 | SULMK[15:8] |  | 0x040 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79
5 | 0x005 | 0x00 | BATMK[7:0] | 5 | 0x005 | 0x00 | BATMK[7:0] |  | 0x050 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95
6 | 0x006 | 0x00 | BATMK[15:8] | 6 | 0x006 | 0x00 | BATMK[15:8] |  | 0x060 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111
7 | 0x007 | 0x00 | ATMSL | 7 | 0x007 | 0x00 | ATMSL |  | 0x070 | 112 | 113 | 114 | 115 | 116 | 117 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104
8 | 0x008 | 0x84 | BGTR | 8 | 0x008 | 0x00 | SPARE |  | 0x080 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 140 | 95 | 96 | 97
9 | 0x009 | 0x60 | TIREF | 9 | 0x009 | 0x00 | SPARE |  | 0x090 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113
10 | 0x00A | 0x0D | REFVM | 10 | 0x00A | 0x01 | COFF_LEDI |  | 0x0A0 | 114 | 115 | 116 | 163 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106
11 | 0x00B | 0xEA | REFIM | 11 | 0x00B | 0x38 | VCVCW |  | 0x0B0 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 186 | 95 | 96 | 97 | 98 | 99
12 | 0x00C | 0x01 | REFUS | 12 | 0x00C | 0x21 | VIPCW |  | 0x0C0 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115
13 | 0x00D | 0x00 | GAINV | 13 | 0x00D | 0x50 | GAINV_POL |  | 0x0D0 | 116 | 209 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108
14 | 0x00E | 0x03 | GAINI | 14 | 0x00E | 0xFF | GINTIOP |  | 0x0E0 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 232 | 95 | 96 | 97 | 98 | 99 | 100 | 101
15 | 0x00F | 0x00 | GAINU | 15 | 0x00F | 0xEF | GAINC |  | 0x0F0 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 255
16 | 0x010 |  | GAINC | 16 | 0x010 | 0xF0 | INTMOP |  | 0x100 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110
17 | 0x011 | 0x00 | Spare | 17 | 0x011 | 0x5D | SLIS |  | 0x110 | 111 | 112 | 113 | 114 | 115 | 116 | 278 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103
18 | 0x012 | 0x00 | Spare | 18 | 0x012 | 0x01 | DFSTL |  | 0x120 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 301 | 95 | 96
19 | 0x013 | 0x00 | Spare | 19 | 0x013 | 0x00 | SPARE |  | 0x130 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112
20 | 0x014 | 0x00 | Spare | 20 | 0x014 | 0x00 | SPARE |  | 0x140 | 113 | 114 | 115 | 116 | 324 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105
21 | 0x015 | 0x00 | Spare | 21 | 0x015 | 0x00 | SPARE |  | 0x150 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 347 | 95 | 96 | 97 | 98
22 | 0x016 | 0x06 | PBLKCW | 22 | 0x016 | 0x06 | PBLKCW |  | 0x160 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114
23 | 0x017 | 0x01 | PBLKPL | 23 | 0x017 | 0x01 | PBLKPL |  | 0x170 | 115 | 116 | 370 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107
24 | 0x018 | 0x01 | PBLKPE | 24 | 0x018 | 0x01 | PBLKPE |  | 0x180 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 393 | 95 | 96 | 97 | 98 | 99 | 100
25 | 0x019 | 0x02 | LPCNTV | 25 | 0x019 | 0xA5 | USEL |  | 0x190 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116
26 | 0x01A | 0x1F | LPCNTCW | 26 | 0x01A | 0x10 | SENPW |  | 0x1A0 | 416 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109
27 | 0x01B | 0x0C | LPCNTIC | 27 | 0x01B | 0x00 | REFUS |  | 0x1B0 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 439 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102
28 | 0x01C | 0x08 | LPCNTPLE | 28 | 0x01C | 0x00 | GAINU |  | 0x1C0 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 462 | 95
29 | 0x01D |  | SEN(0) X | 29 | 0x01D |  | SEN(0) X |  | 0x1D0 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111
30 | 0x01E |  | SEN(1) Y | 30 | 0x01E |  | SEN(1) Y |  | 0x1E0 | 112 | 113 | 114 | 115 | 116 | 485 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104
31 | 0x01F |  | SEN(2) Z | 31 | 0x01F |  | SEN(2) Z |  | 0x1F0 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 508 | 509 | 510 | 511
32 | 0x020 |  | SEN(3) SP6B | 32 | 0x020 |  | SEN(3) SP6B
33 | 0x021 |  | SEN(4) SP5B | 33 | 0x021 |  | SEN(4) SP5B |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: ASIC 1010 EEPROM Default Map
34 | 0x022 |  | SEN(5) SP4B | 34 | 0x022 |  | SEN(5) SP4B |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F
35 | 0x023 |  | SEN(6) SP3B | 35 | 0x023 |  | SEN(6) SP3B |  | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 84 | 60 | 0D | EA | 01 | 00 | 03 | 00
36 | 0x024 |  | SEN(7) SP2B | 36 | 0x024 |  | SEN(7) SP2B |  | 0x010 |  | 00 | 00 | 00 | 00 | 00 | 06 | 01 | 01 | 02 | 1F | 0C | 08
37 | 0x025 |  | SEN(8) SP1B | 37 | 0x025 |  | SEN(8) SP1B |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x020; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 03; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 21; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 07; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 10; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 02; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 20; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 04; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 40
38 | 0x026 |  | SEN(9) CP1B | 38 | 0x026 |  | SEN(9) CP1B |  | 0x030 | A8 | 06 | 2E | 2F | 1B | D3 | 91 | 1B | 08 | 44 | 00 |  | 0F | 00 | 0F | 6C
39 | 0x027 | 0x03 | D1102C10HV | 39 | 0x027 | 0x00 | SENGL[7:0] |  | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 40 | 91 | 00
40 | 0x028 | 0x21 | D1102SD[7:0] | 40 | 0x028 | 0x21 | SNHOP[7:0] |  | 0x050 | C0 | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00
41 | 0x029 | 0x07 | USEL | 41 | 0x029 | 0xBF | GCLKV[7:0] | TEL | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x060; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
42 | 0x02A | 0x10 | SENPW | 42 | 0x02A | 0x88 | GLKV[15:8] |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x070; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 01
43 | 0x02B | 0x02 | SWAIT[7:0] | 43 | 0x02B | 0x02 | SWAIT[15:0] |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x080; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 02
44 | 0x02C | 0x20 | SWAIT[15:8] | 44 | 0x02C | 0x20 | SWAIT[15:0] |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x090; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00
45 | 0x02D | 0x04 | MCOUNT | 45 | 0x02D | 0x20 | MCOUNT |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0A0; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 03; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
46 | 0x02E | 0x40 | DTMSL | 46 | 0x02E | 0xA0 | DTMSL |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0B0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 04
47 | 0x02F |  | PATBL | 47 | 0x02F |  | PATBL |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0C0; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00
48 | 0x030 | 0xA8 | ATOC | 48 | 0x030 | 0x00 | SPARE |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0D0; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 05; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
49 | 0x031 | 0x06 | IOCFG2 | 49 | 0x031 | 0x06 | IOCFG2 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0E0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 06
50 | 0x032 | 0x2E | IOCFG1 | 50 | 0x032 | 0x2E | IOCFG1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0F0; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 07
51 | 0x033 | 0x2F | IOCFG0 | 51 | 0x033 | 0x2F | IOCFG0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x100; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
52 | 0x034 | 0x1B | FIFOTHR | 52 | 0x034 | 0x07 | FIFOTHR |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x110; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 08
53 | 0x035 | 0xD3 | SYNC1 | 53 | 0x035 | 0xD3 | SYNC1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x120; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 09
54 | 0x036 | 0x91 | SYNC0 | 54 | 0x036 | 0x91 | SYNC0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x130; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
55 | 0x037 | 0x1B | PKTLEN | 55 | 0x037 | 0x1B | PKTLEN |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x140; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 0A; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
56 | 0x038 | 0x08 | PKTCTRL1 | 56 | 0x038 | 0x28 | PKTCTRL1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x150; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 0B
57 | 0x039 | 0x44 | PKTCTRL0 | 57 | 0x039 | 0x44 | PKTCTRL0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x160; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00
58 | 0x03A | 0x00 | ADDR | 58 | 0x03A | 0x00 | ADDR |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x170; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 0C; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
59 | 0x03B |  | CHANNR | 59 | 0x03B |  | CHANNR |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x180; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 0D
60 | 0x03C | 0x0F | FSCTRL1 | 60 | 0x03C | 0x0F | FSCTRL1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x190; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00
61 | 0x03D | 0x00 | FSCTRL0 | 61 | 0x03D | 0x00 | FSCTRL0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1A0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 0E; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
62 | 0x03E | 0x0F | FREQ2 | 62 | 0x03E | 0x0F | FREQ2 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1B0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 00; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 0F
63 | 0x03F | 0x6C | FREQ1 | 63 | 0x03F | 0x6C | FREQ1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1C0; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 00; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 00; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 00; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 00; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 00; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 10
64 | 0x040 | 0xCC | FREQ0 | 64 | 0x040 | 0xCC | FREQ0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1D0; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 00; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 00; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 00; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 00; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 00; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
65 | 0x041 | 0x5B | MDMCFG4 | 65 | 0x041 | 0x5B | MDMCFG4 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1E0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 00; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 11
66 | 0x042 | 0xF8 | MDMCFG3 | 66 | 0x042 | 0xF8 | MDMCFG3 |  | 0x1F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 |  |  |  |  | 12 | 00 | 00 | 00
67 | 0x043 | 0x17 | MDMCFG2 | 67 | 0x043 | 0x17 | MDMCFG2
68 | 0x044 | 0x23 | MDMCFG1 | 68 | 0x044 | 0x23 | MDMCFG1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: ASIC 1204 EEPROM Default Map
69 | 0x045 | 0x9A | MDMCFG0 | 69 | 0x045 | 0x9A | MDMCFG0 |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F
70 | 0x046 | 0x47 | DEVIATN | 70 | 0x046 | 0x47 | DEVIATN |  | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 01 | 38 | 21 | 50 | FF | EF
71 | 0x047 | 0x00 | MCSM2 | 71 | 0x047 | 0x00 | MCSM2 |  | 0x010 | F0 | 5D | 01 | 00 | 00 | 00 | 06 | 01 | 01 | A5 | 10 | 00 | 00
72 | 0x048 | 0x30 | MCSM1 | 72 | 0x048 | 0x30 | MCSM1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x020; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 00; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 21; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: BF; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 88; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 02; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 20; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 20; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: A0
73 | 0x049 | 0x18 | MCSM0 | 73 | 0x049 | 0x18 | MCSM0 |  | 0x030 | 00 | 06 | 2E | 2F | 07 | D3 | 91 | 1B | 28 | 44 | 00 |  | 0F | 00 | 0F | 6C
74 | 0x04A | 0x16 | FOCCFG | 74 | 0x04A | 0x16 | FOCCFG |  | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 60 | 91 | 02
75 | 0x04B | 0x6C | BSCFG | 75 | 0x04B | 0x6C | BSCFG |  | 0x050 | 0E | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00
76 | 0x04C | 0x03 | AGCTRL2 | 76 | 0x04C | 0x03 | AGCTRL2 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x060
77 | 0x04D | 0x40 | AGCTRL1 | 77 | 0x04D | 0x60 | AGCTRL1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x070; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 01
78 | 0x04E | 0x91 | AGCTRL0 | 78 | 0x04E | 0x91 | AGCTRL0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x080; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 02
79 | 0x04F | 0x00 | WOREVT1 EVENT0[15:8] | 79 | 0x04F | 0x02 | WOREVT1 EVENT0[15:8] |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x090
80 | 0x050 | 0xC0 | WOREVT0 EVENT0[7:0] | 80 | 0x050 | 0x0E | WOREVT0 EVENT0[7:0] |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0A0; 3 19 35 51 67 83 99 115 108 101 163 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108: 03
81 | 0x051 | 0x78 | WORCTRL | 81 | 0x051 | 0x78 | WORCTRL |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0B0; 10 26 42 58 74 90 106 99 115 108 101 186 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115: 04
82 | 0x052 | 0x56 | FREND1 | 82 | 0x052 | 0x56 | FREND1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0C0
83 | 0x053 | 0x10 | FREND0 | 83 | 0x053 | 0x10 | FREND0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0D0; 1 17 33 49 65 81 97 113 106 99 115 108 101 209 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106: 05
84 | 0x054 | 0xEA | FSCAL3 | 84 | 0x054 | 0xEA | FSCAL3 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0E0; 8 24 40 56 72 88 104 97 113 106 99 115 108 101 232 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113: 06
85 | 0x055 | 0x2A | FSCAL2 | 85 | 0x055 | 0x2A | FSCAL2 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x0F0; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 07
86 | 0x056 | 0x00 | FSCAL1 | 86 | 0x056 | 0x00 | FSCAL1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x100
87 | 0x057 | 0x1F | FSCAL0 | 87 | 0x057 | 0x1F | FSCAL0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x110; 6 22 38 54 70 86 102 95 111 104 97 113 106 99 115 108 101 278 110 103 96 112 105 98 114 107 100 116 109 102 95 111: 08
88 | 0x058 | 0x41 | RCCTRL1 | 88 | 0x058 | 0x41 | RCCTRL1 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x120; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 09
89 | 0x059 | 0x00 | RCCTRL0 | 89 | 0x059 | 0x00 | RCCTRL0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x130
90 | 0x05A | 0x00 | Info Packet PacketID | 90 | 0x05A | 0x00 | Info byte 0 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x140; 4 20 36 52 68 84 100 116 109 102 95 111 104 97 113 106 99 115 108 101 324 110 103 96 112 105 98 114 107 100 116 109: 0A
91 | 0x05B |  | Info byte 1 | 91 | 0x05B |  | Device ID |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x150; 11 27 43 59 75 91 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 347 110 103 96 112 105 98 114 107 100 116: 0B
92 | 0x05C |  | Info byte 2 | 92 | 0x05C |  | Device ID |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x160
93 | 0x05D |  | Info byte 3 | 93 | 0x05D |  | Device ID |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x170; 2 18 34 50 66 82 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 370 110 103 96 112 105 98 114 107: 0C
94 | 0x05E |  | Info byte 4 | 94 | 0x05E |  | Device ID |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x180; 9 25 41 57 73 89 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 393 110 103 96 112 105 98 114: 0D
95 | 0x05F |  | ASIC Identifier | 95 | 0x05F |  | 8X=ASIC 1204, 0X= ASIC 1010 |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x190
96 | 0x060 |  | Year | 96 | 0x060 |  | Year |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1A0; 0 16 32 48 64 80 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 416 110 103 96 112 105: 0E
97 | 0x061 |  | Month | 97 | 0x061 |  | Month |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1B0; 7 23 39 55 71 87 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 439 110 103 96 112: 0F
98 | 0x062 |  | Day | 98 | 0x062 |  | Day |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1C0; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 10
99 | 0x063 |  | Device Type | 99 | 0x063 |  | Device Type |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1D0
100 | 0x064 |  | Device Type | 100 | 0x064 |  | Device Type |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1E0; 5 21 37 53 69 85 101 117 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 485 110: 11
101 | 0x065 |  | Manufacturer / Lot ID | 101 | 0x065 |  | Manufacturer / Lot ID |  | 0x000 0x010 0x020 0x030 0x040 0x050 0x060 0x070 0x080 0x090 0x0A0 0x0B0 0x0C0 0x0D0 0x0E0 0x0F0 0x100 0x110 0x120 0x130 0x140 0x150 0x160 0x170 0x180 0x190 0x1A0 0x1B0 0x1C0 0x1D0 0x1E0 0x1F0: 0x1F0; 12 28 44 60 76 92 108 101 140 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 508: 12; 13 29 45 61 77 93 109 102 95 111 104 97 113 106 99 115 108 101 301 110 103 96 112 105 98 114 107 100 116 109 102 509: 80; 14 30 46 62 78 94 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 97 113 106 99 115 108 101 462 110 103 510: 00; 15 31 47 63 79 95 111 104 97 113 106 99 115 108 101 255 110 103 96 112 105 98 114 107 100 116 109 102 95 111 104 511: 00
102 | 0x066 |  | Manufacturer / Lot ID | 102 | 0x066 |  | Manufacturer / Lot ID
103 | 0x067 |  | Manufacturer / Lot ID | 103 | 0x067 |  | Manufacturer / Lot ID
104 | 0x068 |  | Manufacturer / Lot ID | 104 | 0x068 |  | Manufacturer / Lot ID
105 | 0x069 | 0x00 | Voltage Cal | 105 | 0x069 |  | Accelerometer Cal
106 | 0x06A | 0x00 | Voltage Cal | 106 | 0x06A |  | Accelerometer Cal
107 | 0x06B | 0x00 | Voltage Cal | 107 | 0x06B |  | Accelerometer Cal
108 | 0x06C | 0x00 | Voltage Cal | 108 | 0x06C |  | Accelerometer Cal
109 | 0x06D | 0x00 | Voltage Cal | 109 | 0x06D |  | Accelerometer Cal
110 | 0x06E | 0x00 | Voltage Cal | 110 | 0x06E |  | Accelerometer Cal
111 | 0x06F | 0x00 | Voltage Cal | 111 | 0x06F |  | Voltage H Cal 3.1V
112 | 0x070 | 0x00 | Voltage Cal | 112 | 0x070 |  | Voltage L Cal 3.1V
113 | 0x071 |  | GAINC | 113 | 0x071 |  | Operating Side
114 | 0x072 |  | Cal Factor for CondyleA | 114 | 0x072 |  | Cal Factor for CondyleA
115 | 0x073 |  | Cal Factor for CondyleB | 115 | 0x073 |  | Cal Factor for CondyleB
116 | 0x074 |  | Checksum | 116 | 0x074 |  | Checksum
117 | 0x075 | 0x01 | PacketID 1 | 117 | 0x075 | 0x01 | PacketID 1
140 | 0x08C | 0x02 | PacketID 2 | 140 | 0x08C | 0x02 | PacketID 2
163 | 0x0A3 | 0x03 | PacketID 3 | 163 | 0x0A3 | 0x03 | PacketID 3
186 | 0x0BA | 0x04 | PacketID 4 | 186 | 0x0BA | 0x04 | PacketID 4
209 | 0x0D1 | 0x05 | PacketID 5 | 209 | 0x0D1 | 0x05 | PacketID 5
232 | 0x0E8 | 0x06 | PacketID 6 | 232 | 0x0E8 | 0x06 | PacketID 6
255 | 0x0FF | 0x07 | PacketID 7 | 255 | 0x0FF | 0x07 | PacketID 7
278 | 0x116 | 0x08 | PacketID 8 | 278 | 0x116 | 0x08 | PacketID 8
301 | 0x12D | 0x09 | PacketID 9 | 301 | 0x12D | 0x09 | PacketID 9
324 | 0x144 | 0x0A | PacketID 10 | 324 | 0x144 | 0x0A | PacketID 10
347 | 0x15B | 0x0B | PacketID 11 | 347 | 0x15B | 0x0B | PacketID 11
370 | 0x172 | 0x0C | PacketID 12 | 370 | 0x172 | 0x0C | PacketID 12
393 | 0x189 | 0x0D | PacketID 13 | 393 | 0x189 | 0x0D | PacketID 13
416 | 0x1A0 | 0x0E | PacketID 14 | 416 | 0x1A0 | 0x0E | PacketID 14
439 | 0x1B7 | 0x0F | PacketID 15 | 439 | 0x1B7 | 0x0F | PacketID 15
462 | 0x1CE | 0x10 | PacketID 16 | 462 | 0x1CE | 0x10 | PacketID 16
485 | 0x1E5 | 0x11 | PacketID 17 | 485 | 0x1E5 | 0x11 | PacketID 17
508 | 0x1FC | 0x12 | PacketID 18 | 508 | 0x1FC | 0x12 | PacketID 18
509 | 0x1FD | 0x00 | Spare | 509 | 0x1FD | 0x80 | ID for Echo
510 | 0x1FE | 0x00 | Spare | 510 | 0x1FE | 0x00 | Spare
511 | 0x1FF | 0x00 | Spare | 511 | 0x1FF | 0x00 | Spare

## 1010-Initial
Initial Programming FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | GAINC | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
6003 | 5 |  |  | BF | 84 | TKA | 2016-07-16 | 100 | 100 | 9
1773 | 0A | 00000000 | 00000000 | 06 | 07 | 10 | 00 | 64 | 64 | 09
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID
EEPROM Vector | [10 empty] | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | 2 | 02 | SEN(0) X | [7 empty] | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 0000000000A5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | 2 | 02 | SEN(1) Y | [7 empty] | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 84 | 60 | 0D | EA | 01 | 00 | 03 | 00 | 000000000000000000001B0000000000000084600DEA0100030000
31 | 2 | 02 | SEN(2) Z | [7 empty] | 0x010 | BF | 00 | 00 | 00 | 00 | 00 | 06 | 01 | 01 | 02 | 1F | 0C | 08 | 02 | 02 | 02 | 00000000000000000000BF0000000000060101021F0C0802020200
32 | 4 | 04 | SEN(3) SP6B | [7 empty] | 0x020 | 04 | 04 | 04 | 04 | 04 | 04 | 04 | 03 | 21 | 07 | 10 | 02 | 20 | 04 | 40 | 84 | 000000000000000000000404040404040403210710022004408400
33 | 4 | 04 | SEN(4) SP5B | [7 empty] | 0x030 | A8 | 06 | 2E | 2F | 1B | D3 | 91 | 1B | 08 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 00000000000000000000A8062E2F1BD3911B0844000A0F000F6C00
34 | 4 | 04 | SEN(5) SP4B | [7 empty] | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 40 | 91 | 00 | 00000000000000000000CC5BF817239A47003018166C0340910000
35 | 4 | 04 | SEN(6) SP3B | [7 empty] | 0x050 | C0 | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | 00 | 00 | 00 | 09 | 00000000000000000000C0785610EA2A001F410000000000000900
36 | 4 | 04 | SEN(7) SP2B | [7 empty] | 0x060 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
37 | 4 | 04 | SEN(8) SP1B | [7 empty] | 0x070 | 00 | BF | 64 | 64 | C8 | 01 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 0000000000000000000000BF6464C8010906071017730000000000
38 | 4 | 04 | SEN(9) CP1B | [7 empty] | 0x080 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 02 | 09 | 06 | 07 | 000000000000000000000000000000000000BF6464C80209060700
16 | 191 | BF | GAINC | [7 empty] | 0x090 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 00000000000000000000101773000000000000000000000000BF00
47 | 132 | 84 | PATBL | [7 empty] | 0x0A0 | 64 | 64 | C8 | 03 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000006464C80309060710177300000000000000
59 | 10 | 0A | CHANNR | [7 empty] | 0x0B0 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 04 | 09 | 06 | 07 | 10 | 17 | 00000000000000000000000000000000BF6464C804090607101700
91 | 0 | 00 | Info byte 1 | [7 empty] | 0x0C0 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | 0000000000000000000073000000000000000000000000BF646400
92 | 0 | 00 | Info byte 2 | [7 empty] | 0x0D0 | C8 | 05 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00000000000000000000C805090607101773000000000000000000
93 | 0 | 00 | Info byte 3 | [7 empty] | 0x0E0 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 06 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 0000000000000000000000000000BF6464C8060906071017730000
94 | 0 | 00 | Info byte 4 | [7 empty] | 0x0F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 07 | 000000000000000000000000000000000000000000BF6464C80700
95 | 9 | 09 | ASIC Identifier | [7 empty] | 0x100 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000906071017730000000000000000000000
96 | 6 | 06 | Year | [7 empty] | 0x110 | 00 | 00 | BF | 64 | 64 | C8 | 08 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 000000000000000000000000BF6464C80809060710177300000000
97 | 7 | 07 | Month | [7 empty] | 0x120 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 09 | 09 | 06 | 00000000000000000000000000000000000000BF6464C809090600
98 | 16 | 10 | Day | [7 empty] | 0x130 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000710177300000000000000000000000000
99 | 23 | 17 | Device Type | [7 empty] | 0x140 | BF | 64 | 64 | C8 | 0A | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00000000000000000000BF6464C80A090607101773000000000000
100 | 115 | 73 | Device Type | [7 empty] | 0x150 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 0B | 09 | 06 | 07 | 10 | 0000000000000000000000000000000000BF6464C80B0906071000
101 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x160 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 000000000000000000001773000000000000000000000000BF6400
102 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x170 | 64 | C8 | 0C | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 0000000000000000000064C80C0906071017730000000000000000
103 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x180 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 0D | 09 | 06 | 07 | 10 | 17 | 73 | 000000000000000000000000000000BF6464C80D09060710177300
104 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x190 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 00000000000000000000000000000000000000000000BF6464C800
105 | 0 | 00 | Voltage Cal | [7 empty] | 0x1A0 | 0E | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000E09060710177300000000000000000000
106 | 0 | 00 | Voltage Cal | [7 empty] | 0x1B0 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 0F | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00000000000000000000000000BF6464C80F090607101773000000
107 | 0 | 00 | Voltage Cal | [7 empty] | 0x1C0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 10 | 09 | 0000000000000000000000000000000000000000BF6464C8100900
108 | 0 | 00 | Voltage Cal | [7 empty] | 0x1D0 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
109 | 0 | 00 | Voltage Cal | [7 empty] | 0x1E0 | 00 | BF | 64 | 64 | C8 | 11 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 0000000000000000000000BF6464C8110906071017730000000000
110 | 0 | 00 | Voltage Cal | [7 empty] | 0x1F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | BF | 64 | 64 | C8 | 12 | 00 | 00 | 00 | 000000000000000000000000000000000000BF6464C81200000000
111 | 0 | 00 | Voltage Cal
112 | 0 | 00 | Voltage Cal
113 | 191 | BF | GAINC
114 | 100 | 64 | Cal Factor for CondyleA
115 | 100 | 64 | Cal Factor for CondyleB
116 | 200 | C8 | Checksum

## 1010-GAINC
GAINC Optimization FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | GAINC | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
6003 | 5 | 0 | 0 | BF | 84 | TKA | 2016-07-16 | 100 | 100 | 9
1773 | 0A | 00000000 | 00000000 | 06 | 07 | 10 | 00 | 64 | 64 | 09
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID
EEPROM Vector | [6 empty] | GAINC Calculation Table |  |  |  | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | 2 | 02 | SEN(0) X |  |  |  | Value | Current | Prop. |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 0000000000A5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | 2 | 02 | SEN(1) Y |  |  |  | 2F | 3 | 0.25 |  | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 84 | 60 | 0D | EA | 01 | 00 | 03 | 00 | 000000000000000000001B0000000000000084600DEA0100030000
31 | 2 | 02 | SEN(2) Z |  |  |  | 3F | 4 | 0.33333 |  | 0x010 | #NUM! | 00 | 00 | 00 | 00 | 00 | 06 | 01 | 01 | 02 | 1F | 0C | 08 | 02 | 02 | 02 | #NUM!
32 | 4 | 04 | SEN(3) SP6B |  |  |  | 4F | 5 | 0.41667 |  | 0x020 | 04 | 04 | 04 | 04 | 04 | 04 | 04 | 03 | 21 | 07 | 10 | 02 | 20 | 04 | 40 | 84 | 000000000000000000000404040404040403210710022004408400
33 | 4 | 04 | SEN(4) SP5B |  |  |  | 5F | 6 | 0.5 |  | 0x030 | A8 | 06 | 2E | 2F | 1B | D3 | 91 | 1B | 08 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 00000000000000000000A8062E2F1BD3911B0844000A0F000F6C00
34 | 4 | 04 | SEN(5) SP4B |  |  |  | 6F | 7 | 0.58333 |  | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 40 | 91 | 00 | 00000000000000000000CC5BF817239A47003018166C0340910000
35 | 4 | 04 | SEN(6) SP3B |  |  |  | 7F | 8 | 0.66667 |  | 0x050 | C0 | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | 00 | 00 | 00 | 09 | 00000000000000000000C0785610EA2A001F410000000000000900
36 | 4 | 04 | SEN(7) SP2B |  |  |  | 8F | 9 | 0.75 |  | 0x060 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
37 | 4 | 04 | SEN(8) SP1B |  |  |  | 9F | 10 | 0.83333 |  | 0x070 | 00 | #NUM! | 64 | 64 | #NUM! | 01 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | #NUM!
38 | 4 | 04 | SEN(9) CP1B |  |  |  | AF | 11 | 0.91667 |  | 0x080 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 02 | 09 | 06 | 07 | #NUM!
16 | #NUM! | #NUM! | GAINC |  |  |  | BF | 12 | 1 |  | 0x090 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | #NUM!
47 | 132 | 84 | PATBL |  |  |  | CF | 13 | 1.0833 |  | 0x0A0 | 64 | 64 | #NUM! | 03 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
59 | 10 | 0A | CHANNR |  |  |  | DF | 14 | 1.1667 |  | 0x0B0 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 04 | 09 | 06 | 07 | 10 | 17 | #NUM!
91 | 0 | 00 | Info byte 1 |  |  |  | EF | 15 | 1.25 |  | 0x0C0 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM!
92 | 0 | 00 | Info byte 2 |  |  |  | FF | 16 | 1.3333 |  | 0x0D0 | #NUM! | 05 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
93 | 0 | 00 | Info byte 3 |  |  |  | BF | 12 | 1 |  | 0x0E0 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 06 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | #NUM!
94 | 0 | 00 | Info byte 4 | [7 empty] | 0x0F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 07 | #NUM!
95 | 9 | 09 | ASIC Identifier |  |  |  | Target | 51000 |  |  | 0x100 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000906071017730000000000000000000000
96 | 6 | 06 | Year | [7 empty] | 0x110 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 08 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | #NUM!
97 | 7 | 07 | Month | [7 empty] | 0x120 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 09 | 09 | 06 | #NUM!
98 | 16 | 10 | Day | [7 empty] | 0x130 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000710177300000000000000000000000000
99 | 23 | 17 | Device Type | [7 empty] | 0x140 | #NUM! | 64 | 64 | #NUM! | 0A | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | #NUM!
100 | 115 | 73 | Device Type | [7 empty] | 0x150 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 0B | 09 | 06 | 07 | 10 | #NUM!
101 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x160 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | #NUM!
102 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x170 | 64 | #NUM! | 0C | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
103 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x180 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 0D | 09 | 06 | 07 | 10 | 17 | 73 | #NUM!
104 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x190 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | #NUM!
105 | 0 | 00 | Voltage Cal | [7 empty] | 0x1A0 | 0E | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000E09060710177300000000000000000000
106 | 0 | 00 | Voltage Cal | [7 empty] | 0x1B0 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 0F | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | #NUM!
107 | 0 | 00 | Voltage Cal | [7 empty] | 0x1C0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 10 | 09 | #NUM!
108 | 0 | 00 | Voltage Cal | [7 empty] | 0x1D0 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
109 | 0 | 00 | Voltage Cal | [7 empty] | 0x1E0 | 00 | #NUM! | 64 | 64 | #NUM! | 11 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | #NUM!
110 | 0 | 00 | Voltage Cal | [7 empty] | 0x1F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 64 | 64 | #NUM! | 12 | 00 | 00 | 00 | #NUM!
111 | 0 | 00 | Voltage Cal
112 | 0 | 00 | Voltage Cal
113 | #NUM! | #NUM! | GAINC
114 | 100 | 64 | Cal Factor for CondyleA
115 | 100 | 64 | Cal Factor for CondyleB
116 | #NUM! | #NUM! | Checksum
Counter Output for Zero Load Point for Automatic GAINC Optimization.
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
Battery Failure Test |  |  |  |  | Min L | 18000 | Max L | 24000 |  | Error
Battery | 0 | [5 empty] | PASS? | FALSE |  | 36
Value Maximum Failure Test | [6 empty] | Limit | 65000 |  | Error
Max | 0 | [5 empty] | PASS? | TRUE |  | 0
Anomaly Test | [6 empty] | Limit | 30%
Median | S3 | S4 | S5 | S6 | S7 | S8 | S9 | PASS? |  | Error
#NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! |  | #NUM!
GAINC Optimization
Median | Target | Mult. |  |  |  |  | GAINC | Mult.
#NUM! | 51000 | #NUM! |  |  |  |  | #NUM! | #NUM!
New Values Check
Max L | S3 | S4 | S5 | S6 | S7 | S8 | S9 | PASS? |  | Error
65000 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! |  | #NUM!
Test Result Output to FTS
ErrCodes | #NUM! | GAINC | #NUM!

## 1010-Scales
Initial Programming FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | GAINC | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
6003 | 5 | 0 | 0 | #NUM! | 84 | TKA | 2016-07-16 |  |  | 9
1773 | 0A | 00000000 | 00000000 | 06 | 07 | 10 | 00 |  |  | 09
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID
EEPROM Vector | [11 empty] | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | 2 | 02 | SEN(0) X | [5 empty] | ScaleA | ScaleB |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 0000000000A5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | 2 | 02 | SEN(1) Y | [5 empty] | 0 | 0 |  | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 84 | 60 | 0D | EA | 01 | 00 | 03 | 00 | 000000000000000000001B0000000000000084600DEA0100030000
31 | 2 | 02 | SEN(2) Z |  |  |  | F/Cond [lb] | UnivCal | CalA | CalB |  | 0x010 | #NUM! | 00 | 00 | 00 | 00 | 00 | 06 | 01 | 01 | 02 | 1F | 0C | 08 | 02 | 02 | 02 | #NUM!
32 | 4 | 04 | SEN(3) SP6B |  |  |  | 0 | 0 | 0 | 0 |  | 0x020 | 04 | 04 | 04 | 04 | 04 | 04 | 04 | 03 | 21 | 07 | 10 | 02 | 20 | 04 | 40 | 84 | 000000000000000000000404040404040403210710022004408400
33 | 4 | 04 | SEN(4) SP5B |  |  |  | 5 | 9 | 0 | 0 |  | 0x030 | A8 | 06 | 2E | 2F | 1B | D3 | 91 | 1B | 08 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 00000000000000000000A8062E2F1BD3911B0844000A0F000F6C00
34 | 4 | 04 | SEN(5) SP4B |  |  |  | 10 | 19 | 0 | 0 |  | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 40 | 91 | 00 | 00000000000000000000CC5BF817239A47003018166C0340910000
35 | 4 | 04 | SEN(6) SP3B |  |  |  | 20 | 39 | 0 | 0 |  | 0x050 | C0 | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | 00 | 00 | 00 | 09 | 00000000000000000000C0785610EA2A001F410000000000000900
36 | 4 | 04 | SEN(7) SP2B |  |  |  | 30 | 56 | 0 | 0 |  | 0x060 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
37 | 4 | 04 | SEN(8) SP1B |  |  |  | 40 | 73 | 0 | 0 |  | 0x070 | 00 | #NUM! | 00 | 00 | #NUM! | 01 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | #NUM!
38 | 4 | 04 | SEN(9) CP1B | [8 empty] | 0x080 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 02 | 09 | 06 | 07 | #NUM!
16 | #NUM! | #NUM! | GAINC | [8 empty] | 0x090 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | #NUM!
47 | 132 | 84 | PATBL | [8 empty] | 0x0A0 | 00 | 00 | #NUM! | 03 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
59 | 10 | 0A | CHANNR | [8 empty] | 0x0B0 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 04 | 09 | 06 | 07 | 10 | 17 | #NUM!
91 | 0 | 00 | Info byte 1 | [8 empty] | 0x0C0 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM!
92 | 0 | 00 | Info byte 2 | [8 empty] | 0x0D0 | #NUM! | 05 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
93 | 0 | 00 | Info byte 3 | [8 empty] | 0x0E0 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 06 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | #NUM!
94 | 0 | 00 | Info byte 4 | [8 empty] | 0x0F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 07 | #NUM!
95 | 9 | 09 | ASIC Identifier | [8 empty] | 0x100 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000906071017730000000000000000000000
96 | 6 | 06 | Year | [8 empty] | 0x110 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 08 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | #NUM!
97 | 7 | 07 | Month | [8 empty] | 0x120 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 09 | 09 | 06 | #NUM!
98 | 16 | 10 | Day | [8 empty] | 0x130 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000710177300000000000000000000000000
99 | 23 | 17 | Device Type | [8 empty] | 0x140 | #NUM! | 00 | 00 | #NUM! | 0A | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | #NUM!
100 | 115 | 73 | Device Type | [8 empty] | 0x150 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 0B | 09 | 06 | 07 | 10 | #NUM!
101 | 0 | 00 | Manufacturer / Lot ID | [8 empty] | 0x160 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | #NUM!
102 | 0 | 00 | Manufacturer / Lot ID | [8 empty] | 0x170 | 00 | #NUM! | 0C | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM!
103 | 0 | 00 | Manufacturer / Lot ID | [8 empty] | 0x180 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 0D | 09 | 06 | 07 | 10 | 17 | 73 | #NUM!
104 | 0 | 00 | Manufacturer / Lot ID | [8 empty] | 0x190 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | #NUM!
105 | 0 | 00 | Voltage Cal | [8 empty] | 0x1A0 | 0E | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000E09060710177300000000000000000000
106 | 0 | 00 | Voltage Cal | [8 empty] | 0x1B0 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 0F | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | #NUM!
107 | 0 | 00 | Voltage Cal | [8 empty] | 0x1C0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 10 | 09 | #NUM!
108 | 0 | 00 | Voltage Cal | [8 empty] | 0x1D0 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 000000000000000000000607101773000000000000000000000000
109 | 0 | 00 | Voltage Cal | [8 empty] | 0x1E0 | 00 | #NUM! | 00 | 00 | #NUM! | 11 | 09 | 06 | 07 | 10 | 17 | 73 | 00 | 00 | 00 | 00 | #NUM!
110 | 0 | 00 | Voltage Cal | [8 empty] | 0x1F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | #NUM! | 00 | 00 | #NUM! | 12 | 00 | 00 | 00 | #NUM!
111 | 0 | 00 | Voltage Cal
112 | 0 | 00 | Voltage Cal
113 | #NUM! | #NUM! | GAINC
114 | 0 | 00 | Cal Factor for CondyleA
115 | 0 | 00 | Cal Factor for CondyleB
116 | #NUM! | #NUM! | Checksum
A/D Output for Each Load Point. Populated Automatically during FTS Operation
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
70
140
0
70
140
0
10
20
40
60
80
Battery Failure Test |  |  |  |  | Min L | 18000 | Max L | 24000 |  | Error
BattMax | 0 | BattMin | 0 |  |  |  | PASS? | FALSE |  | 53
Value Maximum/Minimum Failure Test |  |  |  |  | Min L | 40000 | Max L | 65000 |  | Error
ValsMax | 0 | ValsMin | 0 |  |  |  | PASS? | FALSE |  | 63
[11 empty] | 0 | 0
[11 empty] | ScaleA | ScaleB
Compensated, Zero-Baseline Table | [10 empty] | 0 | 0
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | Left | Right |  | F/Cond [lb] | CalA | CalB
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 10 | 0 | 0
40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 20 | 0 | 0
60 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 30 | 0 | 0
80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 40 | 0 | 0
[10 empty] | 80 | 0 | 0
Scale Calculation | [6 empty] | MinScale | 40
Univ | Left | Right | ScaleA | ScaleB |  |  |  | PASS? |  | Error
162 | 0 | 0 | 0 | 0 |  |  |  | FALSE |  | 57
Forces Table
F/Cond [lb] | S3 | S4 | S5 | S6 | S7 | S8 | Left | Right |  | F/Cond [lb] | Left | Right
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 10 | 0 | 0
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 20 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 30 | 0 | 0
40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 40 | 0 | 0
Errors Table | [6 empty] | Limit | 3.5 |  | Error
F/Cond [lb] | Abs/L [lb] | Abs/R [lb] |  |  |  |  | PASS? | FALSE |  | 54
0 | 0 | 0
5 | 5 | 5
10 | 10 | 10
20 | 20 | 20
30 | 30 | 30
40 | 40 | 40
Test Result Output to FTS
ErrCodes | 53 | ScaleA | 0 | ScaleB | 0
Forces Table
F/Cond [lb] | S3 [lb] | S4 [lb] | S5 [lb] | S6 [lb] | S7 [lb] | S8 [lb] | CondA [lb] | CondB [lb] | ErrorA [lb] | ErrorB [lb]
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -5 | -5
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -10 | -10
20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -20 | -20
30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -30 | -30
40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | -40 | -40

## 1204-Initial
Initial Programming FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | Op Side | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
1013 | 5 | 9999 | 13466634 | FF | 84 | TKA | 2016-07-16 | 100 | 100 |  | 1204 EEPROM Programming Prefix
03F5 | 0A | 0000270F | 00CD7C0A | 06 | 07 | 10 | 00 | 64 | 64 | 89 | 13466634 | 12 | 00 | CD | 7C | 0A | 00 | 00 | 00 | 00 | 00
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID | Prefix | 12 | 00 | CD | 7C | 0A | 00 | 00 | 00 | 00 | 00
EEPROM Vector | [10 empty] | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | 2 | 02 | SEN(0) X | [7 empty] | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 1200CD7C0AA5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | 2 | 02 | SEN(1) Y | [7 empty] | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 01 | 38 | 21 | 50 | FF | EF | 1200CD7C0A00000000001B00000000000000000001382150FFEF00
31 | 2 | 02 | SEN(2) Z | [7 empty] | 0x010 | F0 | 5D | 01 | 00 | 00 | 00 | 06 | 01 | 01 | A5 | 10 | 00 | 00 | 02 | 02 | 02 | 1200CD7C0A0000000000F05D01000000060101A510000002020200
32 | 4 | 04 | SEN(3) SP6B | [7 empty] | 0x020 | 04 | 04 | 04 | 04 | 04 | 04 | 04 | 00 | 21 | BF | 88 | 02 | 20 | 20 | A0 | 84 | 1200CD7C0A0000000000040404040404040021BF88022020A08400
33 | 4 | 04 | SEN(4) SP5B | [7 empty] | 0x030 | 00 | 06 | 2E | 2F | 07 | D3 | 91 | 1B | 28 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 1200CD7C0A000000000000062E2F07D3911B2844000A0F000F6C00
34 | 4 | 04 | SEN(5) SP4B | [7 empty] | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 60 | 91 | 02 | 1200CD7C0A0000000000CC5BF817239A47003018166C0360910200
35 | 4 | 04 | SEN(6) SP3B | [7 empty] | 0x050 | 0E | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | CD | 7C | 0A | 89 | 1200CD7C0A00000000000E785610EA2A001F41000000CD7C0A8900
36 | 4 | 04 | SEN(7) SP2B | [7 empty] | 0x060 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 1200CD7C0A000000000006071003F50000270F8000800000000000
37 | 4 | 04 | SEN(8) SP1B | [7 empty] | 0x070 | 00 | FF | 64 | 64 | 11 | 01 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 1200CD7C0A000000000000FF646411018906071003F50000270F00
38 | 4 | 04 | SEN(9) CP1B | [7 empty] | 0x080 | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 02 | 89 | 06 | 07 | 1200CD7C0A00000000008000800000000000FF6464110289060700
15 | 239 | EF | GAINC | [7 empty] | 0x090 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 1200CD7C0A00000000001003F50000270F8000800000000000FF00
47 | 132 | 84 | PATBL |  |  |  |  | Accelerometer Cal Values |  |  | 0x0A0 | 64 | 64 | 11 | 03 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 1200CD7C0A0000000000646411038906071003F50000270F800000
59 | 10 | 0A | CHANNR |  |  |  |  | X | Y |  | 0x0B0 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 04 | 89 | 06 | 07 | 10 | 03 | 1200CD7C0A0000000000800000000000FF64641104890607100300
91 | 0 | 00 | Device ID |  |  |  | Dec | 32768 | 32768 |  | 0x0C0 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 1200CD7C0A0000000000F50000270F8000800000000000FF646400
92 | 205 | CD | Device ID |  |  |  | H | 80 | 80 |  | 0x0D0 | 11 | 05 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 1200CD7C0A000000000011058906071003F50000270F8000800000
93 | 124 | 7C | Device ID |  |  |  | L | 00 | 00 |  | 0x0E0 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 06 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 1200CD7C0A000000000000000000FF646411068906071003F50000
94 | 10 | 0A | Device ID | [7 empty] | 0x0F0 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 07 | 1200CD7C0A000000000000270F8000800000000000FF6464110700
95 | 137 | 89 | 8X=ASIC 1204, 0X= ASIC 1010 |  |  |  |  | Battery Cal Value 3.1V |  |  | 0x100 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 1200CD7C0A00000000008906071003F50000270F80008000000000
96 | 6 | 06 | Year |  |  |  |  | Dec | 0 |  | 0x110 | 00 | 00 | FF | 64 | 64 | 11 | 08 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 1200CD7C0A00000000000000FF646411088906071003F500002700
97 | 7 | 07 | Month |  |  |  |  | H | 00 |  | 0x120 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 09 | 89 | 06 | 1200CD7C0A00000000000F8000800000000000FF64641109890600
98 | 16 | 10 | Day |  |  |  |  | L | 00 |  | 0x130 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A0000000000071003F50000270F800080000000000000
99 | 3 | 03 | Device Type | [7 empty] | 0x140 | FF | 64 | 64 | 11 | 0A | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 1200CD7C0A0000000000FF6464110A8906071003F50000270F8000
100 | 245 | F5 | Device Type |  |  |  |  | Battery Cal Value 2.8V |  |  | 0x150 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0B | 89 | 06 | 07 | 10 | 1200CD7C0A000000000000800000000000FF6464110B8906071000
101 | 0 | 00 | Manufacturer / Lot ID |  |  |  |  | Dec | 0 |  | 0x160 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 1200CD7C0A000000000003F50000270F8000800000000000FF6400
102 | 0 | 00 | Manufacturer / Lot ID |  |  |  |  | H | 00 |  | 0x170 | 64 | 11 | 0C | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 1200CD7C0A000000000064110C8906071003F50000270F80008000
103 | 39 | 27 | Manufacturer / Lot ID |  |  |  |  | L | 00 |  | 0x180 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0D | 89 | 06 | 07 | 10 | 03 | F5 | 1200CD7C0A00000000000000000000FF6464110D8906071003F500
104 | 15 | 0F | Manufacturer / Lot ID | [7 empty] | 0x190 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 1200CD7C0A00000000000000270F8000800000000000FF64641100
105 | 128 | 80 | Accelerometer Cal | [7 empty] | 0x1A0 | 0E | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 1200CD7C0A00000000000E8906071003F50000270F800080000000
106 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1B0 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0F | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 1200CD7C0A0000000000000000FF6464110F8906071003F5000000
107 | 128 | 80 | Accelerometer Cal | [7 empty] | 0x1C0 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 10 | 89 | 1200CD7C0A0000000000270F8000800000000000FF646411108900
108 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1D0 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 1200CD7C0A000000000006071003F50000270F8000800000000000
109 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1E0 | 00 | FF | 64 | 64 | 11 | 11 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 1200CD7C0A000000000000FF646411118906071003F50000270F00
110 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1F0 | 80 | 00 | 80 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 12 | 80 | 00 | 00 | 1200CD7C0A00000000008000800000000000FF6464111280000000
111 | 0 | 00 | Voltage H Cal 3.1V
112 | 0 | 00 | Voltage L Cal 3.1V
113 | 255 | FF | Operating Side
114 | 100 | 64 | Cal Factor for CondyleA
115 | 100 | 64 | Cal Factor for CondyleB
116 | 17 | 11 | Checksum

## 1204-GAINC
Initial Programming FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | Op Side | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
1013 | 5 | 9999 | 13466634 | FF | 84 | TKA | 2016-07-16 | 100 | 100 |  | 1204 EEPROM Programming Prefix
03F5 | 0A | 0000270F | 00CD7C0A | 06 | 07 | 10 | 00 | 64 | 64 | 89 | 13466634 | 12 | 00 | CD | 7C | 0A | 00 | 00 | 00 | 00 | 00
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID | Prefix | 12 | 00 | CD | 7C | 0A | 00 | 00 | 00 | 00 | 00
EEPROM Vector | [10 empty] | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | 2 | 02 | SEN(0) X | [7 empty] | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 1200CD7C0AA5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | 2 | 02 | SEN(1) Y | [7 empty] | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 01 | 38 | 21 | 50 | FF | EF | 1200CD7C0A00000000001B00000000000000000001382150FFEF00
31 | 2 | 02 | SEN(2) Z | [7 empty] | 0x010 | F0 | 5D | 01 | 00 | 00 | 00 | 06 | 01 | 01 | A5 | 10 | 00 | 00 | 02 | 02 | 02 | 1200CD7C0A0000000000F05D01000000060101A510000002020200
32 | 4 | 04 | SEN(3) SP6B | [7 empty] | 0x020 | 04 | 04 | 04 | 04 | 04 | 04 | 04 | 00 | 21 | BF | 88 | 02 | 20 | 20 | A0 | 84 | 1200CD7C0A0000000000040404040404040021BF88022020A08400
33 | 4 | 04 | SEN(4) SP5B | [7 empty] | 0x030 | 00 | 06 | 2E | 2F | 07 | D3 | 91 | 1B | 28 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 1200CD7C0A000000000000062E2F07D3911B2844000A0F000F6C00
34 | 4 | 04 | SEN(5) SP4B | [7 empty] | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 60 | 91 | 02 | 1200CD7C0A0000000000CC5BF817239A47003018166C0360910200
35 | 4 | 04 | SEN(6) SP3B | [7 empty] | 0x050 | 0E | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | CD | 7C | 0A | 89 | 1200CD7C0A00000000000E785610EA2A001F41000000CD7C0A8900
36 | 4 | 04 | SEN(7) SP2B | [7 empty] | 0x060 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A000000000006071003F50000270F0000000000000000
37 | 4 | 04 | SEN(8) SP1B | [7 empty] | 0x070 | 00 | FF | 64 | 64 | 11 | 01 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 1200CD7C0A000000000000FF646411018906071003F50000270F00
38 | 4 | 04 | SEN(9) CP1B | [7 empty] | 0x080 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 02 | 89 | 06 | 07 | 1200CD7C0A00000000000000000000000000FF6464110289060700
15 | 239 | EF | GAINC | [7 empty] | 0x090 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 1200CD7C0A00000000001003F50000270F0000000000000000FF00
47 | 132 | 84 | PATBL | [7 empty] | 0x0A0 | 64 | 64 | 11 | 03 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 1200CD7C0A0000000000646411038906071003F50000270F000000
59 | 10 | 0A | CHANNR | [7 empty] | 0x0B0 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 04 | 89 | 06 | 07 | 10 | 03 | 1200CD7C0A0000000000000000000000FF64641104890607100300
91 | 0 | 00 | Device ID | [7 empty] | 0x0C0 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 1200CD7C0A0000000000F50000270F0000000000000000FF646400
92 | 205 | CD | Device ID | [7 empty] | 0x0D0 | 11 | 05 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 1200CD7C0A000000000011058906071003F50000270F0000000000
93 | 124 | 7C | Device ID | [7 empty] | 0x0E0 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 06 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 1200CD7C0A000000000000000000FF646411068906071003F50000
94 | 10 | 0A | Device ID | [7 empty] | 0x0F0 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 07 | 1200CD7C0A000000000000270F0000000000000000FF6464110700
95 | 137 | 89 | 8X=ASIC 1204, 0X= ASIC 1010 | [7 empty] | 0x100 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A00000000008906071003F50000270F00000000000000
96 | 6 | 06 | Year | [7 empty] | 0x110 | 00 | 00 | FF | 64 | 64 | 11 | 08 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 1200CD7C0A00000000000000FF646411088906071003F500002700
97 | 7 | 07 | Month | [7 empty] | 0x120 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 09 | 89 | 06 | 1200CD7C0A00000000000F0000000000000000FF64641109890600
98 | 16 | 10 | Day | [7 empty] | 0x130 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A0000000000071003F50000270F000000000000000000
99 | 3 | 03 | Device Type | [7 empty] | 0x140 | FF | 64 | 64 | 11 | 0A | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 1200CD7C0A0000000000FF6464110A8906071003F50000270F0000
100 | 245 | F5 | Device Type | [7 empty] | 0x150 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0B | 89 | 06 | 07 | 10 | 1200CD7C0A000000000000000000000000FF6464110B8906071000
101 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x160 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 1200CD7C0A000000000003F50000270F0000000000000000FF6400
102 | 0 | 00 | Manufacturer / Lot ID | [7 empty] | 0x170 | 64 | 11 | 0C | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 1200CD7C0A000000000064110C8906071003F50000270F00000000
103 | 39 | 27 | Manufacturer / Lot ID | [7 empty] | 0x180 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0D | 89 | 06 | 07 | 10 | 03 | F5 | 1200CD7C0A00000000000000000000FF6464110D8906071003F500
104 | 15 | 0F | Manufacturer / Lot ID | [7 empty] | 0x190 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 1200CD7C0A00000000000000270F0000000000000000FF64641100
105 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1A0 | 0E | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A00000000000E8906071003F50000270F000000000000
106 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1B0 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 0F | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 1200CD7C0A0000000000000000FF6464110F8906071003F5000000
107 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1C0 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 10 | 89 | 1200CD7C0A0000000000270F0000000000000000FF646411108900
108 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1D0 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 1200CD7C0A000000000006071003F50000270F0000000000000000
109 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1E0 | 00 | FF | 64 | 64 | 11 | 11 | 89 | 06 | 07 | 10 | 03 | F5 | 00 | 00 | 27 | 0F | 1200CD7C0A000000000000FF646411118906071003F50000270F00
110 | 0 | 00 | Accelerometer Cal | [7 empty] | 0x1F0 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | FF | 64 | 64 | 11 | 12 | 80 | 00 | 00 | 1200CD7C0A00000000000000000000000000FF6464111280000000
111 | 0 | 00 | Voltage H Cal 3.1V
112 | 0 | 00 | Voltage L Cal 3.1V
113 | 255 | FF | Operating Side
114 | 100 | 64 | Cal Factor for CondyleA
115 | 100 | 64 | Cal Factor for CondyleB
116 | 17 | 11 | Checksum
Counter Output for Zero Load Point for Automatic GAINC Optimization.
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
Battery Failure Test |  |  |  |  | Min L | 33000 | Max L | 52000 |  | Error
Battery | 0 | [5 empty] | PASS? | FALSE |  | 35
Value Maximum Failure Test | [6 empty] | Limit | 60000 |  | Error
Max | 0 | [5 empty] | PASS? | TRUE |  | 0
Test Result Output to FTS
ErrCodes | 35 | GAINC | FF | <---Placeholder. No GAINC optimization exists in 1204 yet.

## 1204-Scale-Volt-Accel
Initial Programming FTS Front Panel Contents
DevType | Channel | Lot # | Serial # | Op Side | PATABL | UNI/TKA | Today | ScaleA | ScaleB | Rev
0 | 5 | 0 | 100 | FF | #N/A | #N/A | 2016-07-16 | 100 | 100 |  |  | 1204 EEPROM Programming Prefix
0000 | 0A | 00000000 | 00000064 | 06 | 07 | 10 | 00 | 64 | 64 | 89 |  | 0 | 12 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00
DevType | Channel | Lot # | Serial # | Year | Month | Day | Voltage | ScaleA | ScaleB | ASIC ID |  | Prefix | 12 | 00 | 00 | 00 | 64 | 00 | 00 | 00 | 00 | 00
EEPROM Vector | [11 empty] | EEPROM Array Generator | [16 empty] | EEPROM Copy Cells
29 | #N/A | #N/A | SEN(0) X | [5 empty] | ScaleA | ScaleB |  | 27 Byte | 0x000 | 0x001 | 0x002 | 0x003 | 0x004 | 0x005 | 0x006 | 0x007 | 0x008 | 0x009 | 0x00A | 0x00B | 0x00C | 0x00D | 0x00E | 0x00F | 1200000000A5AAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAAA
30 | #N/A | #N/A | SEN(1) Y | [5 empty] | -150 | -150 |  | 0x000 | 1B | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 00 | 01 | 38 | 21 | 50 | FF | EF | 120000000000000000001B00000000000000000001382150FFEF00
31 | #N/A | #N/A | SEN(2) Z |  |  |  | F/Cond [lb] | UnivCal | CalA | CalB |  | 0x010 | F0 | 5D | 01 | 00 | 00 | 00 | 06 | 01 | 01 | A5 | 10 | 00 | 00 | #N/A | #N/A | #N/A | #N/A
32 | #N/A | #N/A | SEN(3) SP6B |  |  |  | 0 | 0 | 0 | 0 |  | 0x020 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 00 | 21 | BF | 88 | 02 | 20 | 20 | A0 | #N/A | #N/A
33 | #N/A | #N/A | SEN(4) SP5B |  |  |  | 5 | 250 | -375 | -375 |  | 0x030 | 00 | 06 | 2E | 2F | 07 | D3 | 91 | 1B | 28 | 44 | 00 | 0A | 0F | 00 | 0F | 6C | 1200000000000000000000062E2F07D3911B2844000A0F000F6C00
34 | #N/A | #N/A | SEN(5) SP4B |  |  |  | 10 | 500 | -750 | -750 |  | 0x040 | CC | 5B | F8 | 17 | 23 | 9A | 47 | 00 | 30 | 18 | 16 | 6C | 03 | 60 | 91 | 02 | 12000000000000000000CC5BF817239A47003018166C0360910200
35 | #N/A | #N/A | SEN(6) SP3B |  |  |  | 20 | 1000 | -1500 | -1500 |  | 0x050 | 0E | 78 | 56 | 10 | EA | 2A | 00 | 1F | 41 | 00 | 00 | 00 | 00 | 00 | 64 | 89 | 120000000000000000000E785610EA2A001F410000000000648900
36 | #N/A | #N/A | SEN(7) SP2B |  |  |  | 30 | 1500 | -2250 | -2250 |  | 0x060 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 120000000000000000000607100000000000008000800080000000
37 | #N/A | #N/A | SEN(8) SP1B |  |  |  | 40 | 2000 | -3000 | -3000 |  | 0x070 | 00 | FF | 6A | 6A | A2 | 01 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 1200000000000000000000FF6A6AA2018906071000000000000000
38 | #N/A | #N/A | SEN(9) CP1B | [8 empty] | 0x080 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 02 | 89 | 06 | 07 | 120000000000000000008000800080000000FF6A6AA20289060700
15 | 239 | EF | GAINC | [8 empty] | 0x090 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 12000000000000000000100000000000008000800080000000FF00
47 | #N/A | #N/A | PATBL |  |  |  |  | Accelerometer Cal Values |  |  |  | 0x0A0 | 6A | 6A | A2 | 03 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 120000000000000000006A6AA20389060710000000000000800000
59 | 10 | 0A | CHANNR |  |  |  |  | X | Y | Z or 2.8V |  | 0x0B0 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 04 | 89 | 06 | 07 | 10 | 00 | 12000000000000000000800080000000FF6A6AA204890607100000
91 | 0 | 00 | Device ID |  |  |  | Dec | 32768 | 32768 | 32768 |  | 0x0C0 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | 1200000000000000000000000000008000800080000000FF6A6A00
92 | 0 | 00 | Device ID |  |  |  | H | 80 | 80 | 80 |  | 0x0D0 | A2 | 05 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 12000000000000000000A205890607100000000000008000800000
93 | 0 | 00 | Device ID |  |  |  | L | 00 | 00 | 00 |  | 0x0E0 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 06 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 1200000000000000000080000000FF6A6AA2068906071000000000
94 | 100 | 64 | Device ID | [8 empty] | 0x0F0 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 07 | 120000000000000000000000008000800080000000FF6A6AA20700
95 | 137 | 89 | 8X=ASIC 1204, 0X= ASIC 1010 |  |  |  |  | Battery Cal Value 3.1V |  |  |  | 0x100 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 120000000000000000008906071000000000000080008000800000
96 | 6 | 06 | Year |  |  |  |  | Dec | 0 |  |  | 0x110 | 00 | 00 | FF | 6A | 6A | A2 | 08 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 120000000000000000000000FF6A6AA20889060710000000000000
97 | 7 | 07 | Month |  |  |  |  | H | 00 |  |  | 0x120 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 09 | 89 | 06 | 12000000000000000000008000800080000000FF6A6AA209890600
98 | 16 | 10 | Day |  |  |  |  | L | 00 |  |  | 0x130 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | 120000000000000000000710000000000000800080008000000000
99 | 0 | 00 | Device Type | [8 empty] | 0x140 | FF | 6A | 6A | A2 | 0A | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 12000000000000000000FF6A6AA20A890607100000000000008000
100 | 0 | 00 | Device Type |  |  |  |  | Battery Cal Value 2.8V |  |  |  | 0x150 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 0B | 89 | 06 | 07 | 10 | 1200000000000000000000800080000000FF6A6AA20B8906071000
101 | 0 | 00 | Manufacturer / Lot ID |  |  |  |  | Dec | 0 |  |  | 0x160 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 120000000000000000000000000000008000800080000000FF6A00
102 | 0 | 00 | Manufacturer / Lot ID |  |  |  |  | H | 00 |  |  | 0x170 | 6A | A2 | 0C | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 120000000000000000006AA20C8906071000000000000080008000
103 | 0 | 00 | Manufacturer / Lot ID |  |  |  |  | L | 00 |  |  | 0x180 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 0D | 89 | 06 | 07 | 10 | 00 | 00 | 120000000000000000000080000000FF6A6AA20D89060710000000
104 | 0 | 00 | Manufacturer / Lot ID | [8 empty] | 0x190 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 12000000000000000000000000008000800080000000FF6A6AA200
105 | 128 | 80 | Accelerometer Cal | [8 empty] | 0x1A0 | 0E | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 120000000000000000000E89060710000000000000800080008000
106 | 0 | 00 | Accelerometer Cal | [8 empty] | 0x1B0 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 0F | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 12000000000000000000000000FF6A6AA20F890607100000000000
107 | 128 | 80 | Accelerometer Cal | [8 empty] | 0x1C0 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 10 | 89 | 1200000000000000000000008000800080000000FF6A6AA2108900
108 | 0 | 00 | Accelerometer Cal | [8 empty] | 0x1D0 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 120000000000000000000607100000000000008000800080000000
109 | 128 | 80 | Accelerometer Cal | [8 empty] | 0x1E0 | 00 | FF | 6A | 6A | A2 | 11 | 89 | 06 | 07 | 10 | 00 | 00 | 00 | 00 | 00 | 00 | 1200000000000000000000FF6A6AA2118906071000000000000000
110 | 0 | 00 | Accelerometer Cal | [8 empty] | 0x1F0 | 80 | 00 | 80 | 00 | 80 | 00 | 00 | 00 | FF | 6A | 6A | A2 | 12 | 80 | 00 | 00 | 120000000000000000008000800080000000FF6A6AA21280000000
111 | 0 | 00 | Voltage H Cal 3.1V
112 | 0 | 00 | Voltage L Cal 3.1V
113 | 255 | FF | Operating Side
114 | -150 | FFFFFFFF6A | Cal Factor for CondyleA
115 | -150 | FFFFFFFF6A | Cal Factor for CondyleB
116 | 162 | A2 | Checksum
A/D Output for Each Load Point. Populated Automatically during FTS Operation
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
70
140
0
70
140
0 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 20000
10
20
40
60
80
Counts/LB: | -12.5 | -12.5 | -12.5 | -12.5 | -12.5 | -12.5
Battery Failure Test |  |  |  |  | Min L | 33000 | Max L | 52000 |  | Error
BattMax | 20000 | BattMin | 20000 |  |  |  | PASS? | FALSE |  | 52
Count Min/Max Test |  |  |  |  | Min L | 5000 | Max L | 60000 |  | Error
CntMax | 1000 | CntMin | 1000 |  |  |  | PASS? | FALSE |  | 63
[11 empty] | -75 | -75
[11 empty] | ScaleA | ScaleB
Zero-Baseline Table | [10 empty] | -150 | -150 |  | New loads converted
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | Left | Right |  | F/Cond [lb] | CalA | CalB |  | LB Left | LB Right | SUM | PASS? |  |  |  |  | Increasing loads?
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 |  | S3 | S4 | S5 | S6 | S7 | S8 | LIMIT
10 | -1000 | -1000 | -1000 | -1000 | -1000 | -1000 | -3000 | -3000 |  | 5 | -375 | -375 |  | 40 | 40 | 80 | 62 |  | 62 | 62 | 62 | 62 | 62 | 62 | 4
20 | -1000 | -1000 | -1000 | -1000 | -1000 | -1000 | -3000 | -3000 |  | 10 | -750 | -750 |  | 40 | 40 | 80 | 62 |  | 62 | 62 | 62 | 62 | 62 | 62 | 9
40 | -1000 | -1000 | -1000 | -1000 | -1000 | -1000 | -3000 | -3000 |  | 20 | -1500 | -1500 |  | 40 | 40 | 80 | 62 |  | 62 | 62 | 62 | 62 | 62 | 62 | 19
60 | -1000 | -1000 | -1000 | -1000 | -1000 | -1000 | -3000 | -3000 |  | 30 | -2250 | -2250 |  | 40 | 40 | 80 | 62 |  | 62 | 62 | 62 | 62 | 62 | 62 | 29
80 | -1000 | -1000 | -1000 | -1000 | -1000 | -1000 | -3000 | -3000 |  | 40 | -3000 | -3000 |  | 40 | 40 | 80 | 0 |  | 62 | 62 | 62 | 62 | 62 | 62 | 39
[10 empty] | 80 | -6000 | -6000 |  |  |  |  | Error |  | Error | Error | Error | Error | Error | Error
NEW Scale Calculation |  |  |  |  | MaxScale | 70 | MinScale | 30 | [8 empty] | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
Univ | Left | Right | ScaleA | ScaleB |  |  |  | PASS? |  | Error | [6 empty] | 62 |  | 66 | 67 | 68 | 69 | 70 | 71
2000 | -3000 | -3000 | -150 | -150 |  |  |  | FALSE |  | 57 | [10 empty] | MIN counts at 80lb
|  |  | -0.013333 | -0.013333 | [14 empty] | S3 | S4 | S5 | S6 | S7 | S8
Forces Table | [18 empty] | Error | Error | Error | Error | Error | Error
F/Cond [lb] | S3 | S4 | S5 | S6 | S7 | S8 | Left | Right |  | F/Cond [lb] | Left | Right | [6 empty] | 72 | 73 | 74 | 75 | 76 | 77
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0
5 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 |  | 5 | 40 | 40
10 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 |  | 10 | 40 | 40
20 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 |  | 20 | 40 | 40
30 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 |  | 30 | 40 | 40
40 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 |  | 40 | 40 | 40
Errors Table | [6 empty] | Limit | 3.5 |  | Error
F/Cond [lb] | Abs/L [lb] | Abs/R [lb] | Pre/L | Pre/R |  |  | PASS? | FALSE |  | 54
0 | 0 | 0
5 | 35 | 35 | 19 | 19 | Avg | -2 | -
10 | 30 | 30 | 14.3178 | 14.3178 | StDev | 1 | 2
20 | 20 | 20 | 1 | 1 | Slope | 0.05 | -
30 | 10 | 10 | TRUE | TRUE | Nonlinearity? |  |  |  |  | Prog Error
40 | 0 | 0 | FALSE | FALSE | Preload? |  |  |  |  | 0
Test Result Output to FTS
ErrCodes | 77 | ScaleA | -150 | ScaleB | -150
Forces Table
F/Cond [lb] | S3 [lb] | S4 [lb] | S5 [lb] | S6 [lb] | S7 [lb] | S8 [lb] | CondA [lb] | CondB [lb] | ErrorA [lb] | ErrorB [lb]
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 35 | 35
10 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 30 | 30
20 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 20 | 20
30 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 10 | 10
40 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 13.3333 | 40 | 40 | 0 | 0
A/D Output for Each Alignment Point. Populated by Operator by Orienting and Clicking
Orientation | [5 empty] | AccelX | AccelY | AccelZ
VERASENSE is Oriented FLAT on Level Surface | [5 empty] | 32768 | 32768 | 58000
VERASENSE is Oriented at 45 DEGREES, Posterior Down | [5 empty] | 48000 | 32768 | 48000
VERASENSE is Oriented VERTICALLY, Posterior Down | [5 empty] | 58000 | 48000 | 32768
Accelerometer PASS/FAIL | [6 empty] | BMT? | FALSE
1g, Max | 65000 | 1g, Min | 45000 |  | 0g, Max | 42000 | 0g, Min | 22000
Accelerometer Channel |  |  |  |  | 1g | PASS? | 0g | PASS? |  | Error | Error
Accelerometer, X |  |  |  |  | 58000 | TRUE | 32768 | TRUE |  | 0 | 0
Accelerometer, Y |  |  |  |  | N/A | TRUE | 32768 | TRUE |  | 0 | 0
Accelerometer, Z |  |  |  |  | 58000 | TRUE | 32768 | TRUE |  | 0 | 0
[9 empty] | Values read from the device EEPROM | [6 empty] | Min | Max
Final | [8 empty] | 2.8V | 3.1V | ScaleA | ScaleB | V | 3 | PCB sensor and Batt | 8000 | 15000 | There are 4000 counts/0.1V. 10000 was to high, so 8000 was chosen.
Battery Voltage | [7 empty] | 0 | 0 | 0 | 0 | 0 | 0 | Batt last |  | Batt Diff? | Batt Diff?
[10 empty] | Error |  |  |  | Error | Error | Error | Error
[10 empty] | 52 |  |  |  | 52 | 43 | 43
Load Accuracy Test. A/D Output for Each Load Stable Point. (no yet used)
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
20
40
80
Load Accuracy Test. A/D Output after Damping Filter. (no yet used)
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
20
40
80
Load Accuracy Test. A/D Output after Averaging Filter. (no yet used)
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
20
40
80
Load Accuracy Test. A/D Output after Averaging the Filtered Data. (no yet used)
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | S9 | Battery
0
20
40
80
Load Accuracy Test. Forces Table
Force [lb] | S3 | S4 | S5 | S6 | S7 | S8 | Left | Right |  | Error | Error |  | Correction Factor
0 | [6 empty] | 0 | 0 |  | 0 | 0
20 | [6 empty] | 0 | 0 |  | 80 | 81
40 | [6 empty] | 0 | 0 |  | 82 | 83
80 | [6 empty] | 0 | 0 |  | 84 | 85
Load Accuracy Test. Errors Table | [6 empty] | Limit | 3.5 |  | Error
F/Cond [lb] | Abs/L [lb] | Abs/R [lb] | Pre/L | Pre/R |  |  | PASS? | FALSE |  | 0
0 | 0 | 0
10 | 10 | 10 | 23.3333 | 23.3333 | Avg | -2 | -
20 | 20 | 20 | 15.2753 | 15.2753 | StDev | 1 | 2
40 | 40 | 40 | 1 | 1 | Slope | 0.05 | -
|  |  | TRUE | TRUE | Nonlinearity?
|  |  | FALSE | FALSE | Preload?